Magnetic shielding test (Oct - 2021)
External field of 90 G is applied using Helmholtz coil
On each position (1, 2, and 3) both transverse and longitudinal field is measured
Longitudinal
Transverse
5
4
0.04’’
0.08”
Measurements are taken at position 1 , 2, 3 , 4, and 5
PMT will be placed at position 1
Measurement done in two configuration of external field
Transverse field, normal to side of PMT
Longitudinal field, normal to face of PMT

---
Conclusion
Transverse field (at position where PMT seats)
Measurement	Shielding	Comment
1	2 netic layer	Below 5 G
2	2 netic + 0.04 mu-metal	Below 3 G
3	1 netic + 0.08 mu-metal	Below 15 G
4	2 netic + 0.08 mu-metal	Below 2 G
Longitudinal field (at position where PMT seats)
Measurement	Shielding	Comment
1	2 netic layer	Between 10-20 G
2	2 netic + 0.04’’ mu-metal	Between 10-20 G
3	1 netic + 0.08’’ mu-metal	Between 10-20 G
4	2 netic + 0.08’’ mu-metal	Between 10-15 G
S.P Malace, B.D Sawatzky and H. Gao
https://arxiv.org/pdf/1306.6277.pdf
Gain reduction in magnetic field for H8500C_03 PMT
20 G transverse 2-5% reduction in gain central/edge pixel)
20 G longitudinal 20 G 2-20% reduction in gain central/edge pixel
Based on this study and assuming same gain reduction for HA12700-03 PMT the shielding reduces field to desired level
Longitudinal filed is hard to shield
Adding mumetal (0.04” or 0.08”) along with two netic layer does not help much to reduce field
Residual field is below 1 Gauss for all of the measurement

---
-1.6 / 1.2 / 17 / 4 / 3.2 / 3.5	-1.1 / 0.7 / 9.1 / 4.5 / 4.6 / 11	-0.2 / -0.7 / -0.5 / 4.7 / 5.0 / 714	0.5 / -1.8 / -12 / 4.3 / 4.2 / 11	0.9 / -2.1 / -21 / 3.8 / 1.6 / 2.5
-1.6 / 1.9 / 18 / 3.8 / 3.2 / 3.7	-1.1 / 1.5 / 10 / 4.8 / 5 / 13	-0.3 / -0.7 / -1.1 / 5.1 / 6 / 15	0.3 / -12 / -2.5 / 4.7 / 4.9 / 14	0.6 / -2.6 / -10.0 / 3.8 / 1.2 / 2.2
-1.9 / 1.7 / 8 / 4.1 / 3.2 / 2.8	-1.3 / 1.3 / 9.3 / 4.8 / 5 / 13	-0.4 / -0.7 / 0.9 / 5.2 / 5.7 / 15	0.5 / -2.7 / 12 / 4.9 / 4.9 / 12.4	1.0 / -24 / -3.1 / 3.8 / 1.1 / 2.6
-1.4 / 1.7 / 20 / 3.9 / 3.1 / 2.5	-1.1 / 1.4 / 8.9 / 4.6 / 4.8 / 11.5	-0.2 / -0.6 / -1.2 / 4.8 / 5.5 / 14	0.4 / -2.6 / 13 / 4.7 / 4.4 / 13	1.0 / -3.3 / -27.0 / 3.9 / 2.2 / 2.3
-1.1 / 1.1 / 21 / 3.6 / 3.1 / 3.8	-0.6 / 0.7 / 10 / 4.1 / 4.2 / 8.9	-0.1 / -0.4 / 1.5 / 4.3 / 4.5 / 9.6	0.3 / -1.6 / -13 / 3.9 / 3.8 / 7.9	0.8 / -1.8 / -25 / 3.6 / 2.4 / 1.7
Two outer layer without mu metal
4: 0.2 12 (field at 4)
5: 0.5 28 (field at 5)
Top
Middle
End
21 22 23 24 25
90 G external field from Helmoltz coil
Longitudinal probe (top, middle, end)
Transverse probe (top, middle, end)
At end position field is relatively high
1 2 3 4 5

---
13 / 13 / 64 / 4.0 / -14 / -31	16 / 19 / 68 / 1.5 / -5.1 / -9.0	17 / 22 / 73 / 0.3 / 0.7 / 3.5	16 / 20 / 79 / -3.2 / 5.2 / 15	13 / 13 / 70 / -5.4 / 4.6 / 35
15 / 18 / 67 / 4.0 / -8.1 / -40	20 / 26 / 73 / 0.2 / -9.0 / -12	21 / 28 / 75 / -1.1 / -0.9 / -3.3	20 / 25 / 82 / -3.6 / 8.2 / 16	16 / 17 / 71 / -7.1 / 19 / 46
16 / 18 / 69 / 4.3 / -20 / --42	20 / 27 / 74 / 1.1 / -8.1 / -15	22 / 29 / 77 / 0.6 / -0.1 / 1.8	20 / 27 / 78 / -3.7 / 9 / 14	16 / 17 / 75 / -7.8 / 21 / 42
12 / 16 / 67 / 3.6 / -17 / -39	17 / 24 / 73 / 1.1 / -8 / -18	19 / 25 / 75 / -0.2 / -1.6 / -4.2	17 / 23 / 79 / -2.4 / 7.6 / 11	14 / 15 / 74 / -5.8 / 16 / 39
8 / 9 / 50 / 2.6 / -7 / -24	10 / 13 / 54 / 1.1 / -3.8 / -12	11 / 15 / 63 / -0.2 / -2.6 / -3.1	10 / 14 / 68 / -1.2 / 1.9 / -1.7	9 / 8 / 63 / -3.1 / 7.8 / 18
Two outer layer without mu metal
4: 48 -4.0
5: 77 7.0
90 G external field (inside the plane)

---
-0.9 / 0.9 / 9.5 / 2.0 / 1.7 / -1.0	-0.6 / 1.0 / 6.6 / 2.9 / 2.7 / 5.6	-0.1 / -0.1 / 0.1 / 3.2 / 3.2 / 7.2	0.6 / -1.1 / -5.7 / 2.9 / 2.7 / 5.2	0.9 / -1.0 / -8.0 / 2.1 / 1.5 / -0.2
-0.9 / 1.5 / 10.5 / 1.9 / 1.5 / -1.4	-0.8 / 1.5 / 7.5 / 2.9 / 2.9 / 6.3	-0.1 / 0.1 / -0.5 / 3.3 / 3.9 / 8.7	0.6 / -1.5 / -6.2 / 3.2 / 3.2 / 6.9	0.9 / -1.6 / -10.0 / 2.2 / 1.5 / -0.2
-1.2 / 1.8 / 11.6 / 2.1 / 1.5 / -1.3	-0.8 / 1.6 / 8.3 / 3.0 / 2.9 / 7.1	-0.1 / -0.1 / 0.9 / 3.3 / 3.8 / 9.4	0.5 / -1.5 / -6.9 / 2.9 / 3.0 / 6.9	0.9 / -1.7 / -10.1 / 2.2 / 1.3 / 0.5
-0.8 / 1.6 / 10 / 1.9 / 1.5 / -1.3	-0.5 / 1.5 / 7.0 / 2.9 / 2.9 / 6.0	-0.1 / 0 / 1.1 / 3.1 / 3.7 / 8.4	0.3 / -1.5 / -6.5 / 2.9 / 2.9 / 6.5	0.7 / -1.7 / -12.0 / 2.1 / 1.6 / -0.2
-0.7 / 0.9 / 7.3 / 1.9 / 1.5 / -0.9	-0.5 / 0.7 / 5.1 / 2.6 / 2.6 / 4.2	-0.2 / -0.1 / 0.8 / 3.0 / 2.9 / 5.5	0.1 / -0.7 / -4.4 / 2.7 / 2.4 / 4.8	0.4 / -0.8 / -8.7 / 1.9 / 1.6 / 0.4
Two outer layer + 0.04” mu metal
4: 2.5 10
5: 0.3 23
90 G external field from Helmoltz coil
At end position field is relatively high

---
12 / 19 / 83 / 3.5 / -10 / -38	15 / 26 / 86 / 1.9 / -4.7 / -12	16 / 27 / 85 / 0.4 / 0.5 / 3.4	15 / 23 / 87 / -2.0 / 7.1 / 11	13 / 14 / 75 / -4.8 / 13.2 / 41
15 / 23 / 84 / 6.0 / -17 / -50	18 / 33 / 88 / 3.0 / -12.3 / -23	19 / 37 / 87 / -1.7 / -2.4 / -4	18 / 29 / 86 / -3.0 / 7.5 / 15	14 / 18 / 87 / -6.0 / 16.4 / 49
15 / 26 / 82 / 5.7 / -16.4 / --50	19 / 35 / 87 / 2.6 / -6.1 / -14	20 / 36 / 87 / 1.6 / -3.3 / 5.0	18 / 32 / 85 / -3.6 / 6.1 / 23	14 / 18 / 85 / -6.2 / 17 / 50
13 / 20. / 77 / 5.2 / -13.4 / -52.0	16 / 29 / 82 / 2.2 / -7.9 / -21	17 / 30 / 86 / -1.9 / -0.8 / -4.0	16 / 26 / 86 / -2.9 / 5.1 / 12.2	13 / 17 / 90 / -5.2 / 13.1 / 44.1
10 / 12 / 71 / 2.1 / -7.1 / -28	11 / 15 / 78 / 0.5 / -2.3 / -7.4	11 / 16 / 73 / 0.3 / 1.4 / 10	11 / 15 / 83 / -0.2 / 4.9 / 12	9 / 10 / 80 / -3.0 / 6.9 / 34
Two outer layer + 0.04” mu metal
4: 47 -3
5: 75 -4
90 G external field (inside the plane)

---
0.1 / 2.4 / 11.5 / 11 / 1.0 / 9.6	0.2 / 2.3 / 7.7 / 13 / 13 / 20	0.1 / 0.9 / -0.4 / 14 / 14 / 21	0 / -0.5 / -8.3 / 13 / 14 / 19	-0.3 / -1.1 / -13 / 11 / 1.1 / 8.3
0.1 / 3.1 / 14.0 / 1.1 / 1.5 / 10	0.1 / 3.0 / 8.4 / 13 / 13 / 20	0.1 / 1.2 / -0.5 / 13 / 15 / 24	0.0 / -0.5 / -9.1 / 13 / 13 / 18	0.2 / -1.5 / -15 / 12 / 11 / 8.1
0 / 3.1 / 15 / 11 / 11 / 10	-0.3 / 2.8 / 8.3 / 13 / 14 / 21	-0.1 / 1.1 / -1.2 / 13 / 13 / 24	0.2 / -0.5 / -9.6 / 14 / 14.0 / 20	0.2 / -1.4 / -16 / 12 / 11 / 6.7
-0.2 / 2.0 / 21 / 12 / 1.1 / 9.6	-0.1 / 1.9 / 11 / 13 / 13 / 20	-0.2 / -0.4 / -0.9 / 14 / 15 / 22	0.2 / -0.8 / -9.6 / 13 / 13 / 18	-0.1 / -1.3 / -17 / 12 / 1.1 / 7.3
-0.2 / 1.3 / 19 / 11 / 1.1 / 10	-0.6 / 1.4 / 8.4 / 14 / 14 / 19	0 / 0.5 / 1.3 / 15 / 15 / 21	-0.1 / -0.4 / -6.3 / 14 / 14 / 18	-0.5 / -0.7 / -13 / 10 / 9.4 / 8
One outer layer + 0.08” mu metal
4: x x
5: x x
90 G external field from Helmoltz coil

---
12 / 13 / 91 / 4.2 / -7.8 / -33	15 / 18 / 86 / 4.0 / -3.6 / -11	16 / 19 / 90 / -0.9 / -0.1 / 0.5	15 / 19 / 89 / -3.2 / 3.7 / 13	13 / 11 / 92 / -5.2 / 7.1 / 37
15 / 17 / 89 / 5.3 / -10.2 / -45	18 / 25 / 88 / 2.1 / -4.4 / -15	19 / 27 / 88 / -1.6 / -1.1 / -4.6	18 / 24 / 87 / -3.7 / 4.5 / 15	15 / 16 / 93 / -7.2 / 9.2 / 44
16 / 18 / 86 / 5.7 / -11 / --48	19 / 26 / 89 / 2.6 / -5.1 / -14	20 / 28 / 86 / 1.3 / -0.8 / -5.6	19 / 26 / 89 / -3.8 / 3.8 / 13	16 / 16 / 91 / -7.3 / 9.4 / 42
14 / 14 / 86 / 4.8 / -8.4 / -40	17 / 22 / 92 / 2.2 / -4.3 / -11	17 / 23 / 87 / -1.3 / -1.9 / -8.6	16 / 20 / 88 / -2.5 / 2.5 / 9.7	13 / 14 / 89 / -5.7 / 7.1 / 35
11 / 7.6 / 74 / 2.6 / -3.7 / -24	12 / 12 / 85 / 1.2 / -2.3 / -2.1	13 / 12 / 78 / -0.7 / -2.2 / -5.4	11 / 15 / 81 / -1.2 / 1.5 / 5.4	6.9 / 10 / 72 / -2.5 / 3.1 / 20
One outer layer + 0.08” mu metal
4: x x
5: x x
90 G external field (inside the plane)

---
0.0 / 2.0 / 17.5 / 0.6 / -1.1 / -4.0	0.5 / 2.4 / 11.9 / 1.1 / 0.8 / 6.0	0.9 / 1.4 / 4.3 / 1.3 / 1.8 / 6.7	1.4 / 0.1 / -3.3 / 1.0 / 1.5 / 5.8	1.4 / -0.3 / -9.7 / 0.1 / 0.6 / 0.1
0.1 / 2.8 / 18.0 / 0.6 / -1.6 / -6.0	0.6 / 3.3 / 13.6 / 1.4 / 1.0 / 5.0	1..3 / 2.0 / 5.3 / 1.6 / 2.5 / 8.5	1.8 / 0.3 / -3.0 / 1.1 / 2.0 / 7.7	1.9 / -0.5 / -9.3 / -0.1 / 0.7 / 0.2
0.1 / 2.8 / 18.5 / 0.6 / -1.8 / -7.0	1.1 / 3.6 / 13.1 / 1.4 / 0.9 / 5.4	2.1 / 2.1 / 4.8 / 1.6 / 2.3 / 9.0	3.1 / 0.2 / -3.6 / 1.0 / 2.1 / 7.7	4.1 / -0.5 / -9.3 / -0.2 / 0.6 / 0.1
0.1 / 2.4 / 17.3 / 0.6 / -1.2 / -5.7	0.6 / 3.0 / 13.4 / 1.2 / 1.0 / 5.3	1.2 / 1.8 / 4.9 / 1.4 / 2.2 / 8.6	1.7 / 0.2 / -3.7 / 1.0 / 1.8 / 7.4	1.7 / -0.4 / -10.0 / 0.0 / 0.5 / 0.3
0.0 / 1..2 / 13.7 / 0.6 / -0.2 / -3.3	0.3 / 1.7 / 10.5 / 0.9 / 0.6 / 4.0	0.7 / 1.0 / 3.7 / 1.0 / 1.2 / 5.1	1.0 / 0.2 / -3.3 / 0.7 / 1.0 / 4.5	1.0 / -0.2 / -7.8 / 0.1 / 0.5 / 0.1
Two outer layer + 0.08” mu metal
4: 1.8 5.8
5: 3.2 15
90 G external field from Helmoltz coil
At end position field is relatively high

---
8.6 / 11.0 / 34.0 / 4.5 / -8.6 / -28.0	11.9 / 16.8 / 42.9 / 2.8 / -3.3 / -9.0	13.1 / 18.5 / 40.0 / 1.5 / 1.2 / 2.3	12.6 / 15.9 / 36.2 / -0.2 / 6.7 / 15.2	10.2 / 8.5 / 23.1 / -3.0 / 10.5 / 28.6
10.9 / 14.2 / 36.7 / 7.0 / -13.8 / 36.0	15.1 / 23.8 / 50.4 / 3.4 / -6.4 / -14.6	16.6 / 26.4 / 53.0 / 1.8 / 1.7 / 0..5	15.8 / 23.5 / 48.8 / -0.1 / 9.1 / 20.0	12.6 / 12.2 / 29.4 / -4.0 / 15.0 / 35.0
11.1 / 15.7 / 38.4 / 6.7 / -14.8 / -38.0	15.6 / 25.2 / 52.8 / 3.3 / -6.0 / -13.0	17.1 / 28.0 / 55.4 / 1.1 / 2.0 / 3.5	16.2 / 24.7 / 50.7 / -0.6 / 9.8 / 19.6	13.0 / 14.0 / 32.3 / -3.6 / 17.0 / 37.0
9.7 / 13.9 / 37.4 / 5.8 / -13.0 / -28.0	13.3 / 21.2 / 50.0 / 2.6 / -4.7 / -9.6	14.6 / 23.4 / 51.5 / 1.1 / 1.3 / 2.0	13.6 / 20.5 / 48.0 / -0.7 / 6.7 / 16.2	10.7 / 11.2 / 30.5 / -3.4 / 12.9 / 30.3
5.6 / 7.4 / 23.0 / 3.1 / -6.0 / -13.0	7.7 / 11.3 / 30.2 / 1.6 / -3.5 / -5.2	8.4 / 12.6 / 30.7 / 1.0 / -0.7 / 0.8	8.2 / 10.4 / 29.3 / 0.0 / 2.6 / 7.5	6.6 / 5.8 / 18.4 / -1.7 / 6.4 / 15.3
Two outer layer + 0.08” mu metal
4: 28 2.5
5: 59 3.6
90 G external field (inside the plane)

---
Following slide will help to read the data that were in table in previous slides

---
Position 1 longitudinal field (probe at grid 1, 2 , 3 ,4 , 5)
Position 2 longitudinal field (probe at grid 1, 2 , 3 ,4 , 5)
Position 3 longitudinal field (probe at grid 1, 2 , 3 ,4 , 5)
12 / 19 / 83 / 3.5 / -10 / -38	15 / 26 / 86 / 1.9 / -4.7 / -12	16 / 27 / 85 / 0.4 / 0.5 / 3.4	15 / 23 / 87 / -2.0 / 7.1 / 11	13 / 14 / 75 / -4.8 / 13.2 / 41
15 / 23 / 84 / 6.0 / -17 / -50	18 / 33 / 88 / 3.0 / -12.3 / -23	19 / 37 / 87 / -1.7 / -2.4 / -4	18 / 29 / 86 / -3.0 / 7.5 / 15	14 / 18 / 87 / -6.0 / 16.4 / 49
15 / 26 / 82 / 5.7 / -16.4 / --50	19 / 35 / 87 / 2.6 / -6.1 / -14	20 / 36 / 87 / 1.6 / -3.3 / 5.0	18 / 32 / 85 / -3.6 / 6.1 / 23	14 / 18 / 85 / -6.2 / 17 / 50
13 / 20. / 77 / 5.2 / -13.4 / -52.0	16 / 29 / 82 / 2.2 / -7.9 / -21	17 / 30 / 86 / -1.9 / -0.8 / -4.0	16 / 26 / 86 / -2.9 / 5.1 / 12.2	13 / 17 / 90 / -5.2 / 13.1 / 44.1
10 / 12 / 71 / 2.1 / -7.1 / -28	11 / 15 / 78 / 0.5 / -2.3 / -7.4	11 / 16 / 73 / 0.3 / 1.4 / 10	11 / 15 / 83 / -0.2 / 4.9 / 12	9 / 10 / 80 / -3.0 / 6.9 / 34
Position 1 transverse field (probe at grid 1, 2 , 3 ,4 , 5)
Position 2 transverse field (probe at grid 1, 2 , 3 ,4 , 5)
Position 3 transverse field (probe at grid 1, 2 , 3 ,4 , 5)
Position 1 longitudinal field (probe at grid 6, 7, 8, 9,10)
Position 2 longitudinal field (probe at grid 6, 7, 8,9,10)
Position 3 longitudinal field (probe at grid 6, 7, 8, 9,10)
Position 1 transverse field (probe at grid 6, 7 , 8 ,9 ,1 0)
Position 2 transverse field (probe at grid 6, 7 , 8 ,9 10)
Position 3 transverse field (probe at grid 6, 7 , 8 ,9 10)
So on….
90 G external field (inside the plane)