## Slide 1: GEM chambers for SoLID Nilanga Liyanage University of Virginia

## Slide 2: Tracking needs for SoLID (PVDIS)
Rate: from 100 kHz to 600 kHz (with baffles), GEANT3 estimation
Spatial Resolution: 0.2 mm (sigma)
Total area: 23 m2 total area (30 sectors x4 planes, each sector cover 10 degree)
Need to be Magnetic field tolerant
Lumi = 5.4E38 /cm2/s

## Slide 3: Choice of the technology
System Requirements | Tracking Technology |  | 
 | Drift | MPGD | Silicon
High Rate (up to): 1 MHz/cm2 | NO | MHz/mm2 | MHz/mm2
Resolution: 200 mm | Achievable | 50 mm | 30 mm
Large Area: Single chamber up to 0.5 m2 total 23 m2 | YES | Doable | Very Expensive
Flexibility in readout geometry and lower spark rate
GEM
mMs
New recent development in resistive mM and GEM + mM to improve spark rate

## Slide 4: GEM working principle
Ionization
Multiplication
(x20)
Readout
Multiplication
(x20)
Multiplication
(x20)
Recent technology: F. Sauli, Nucl. Instrum. Methods A386(1997)531
GEM foil: 50 mm Kapton + few mm copper on both sides with 70 mm holes, 140 mm pitch
Strong electrostatic field in the GEM holes

## Slide 5
COMPASS readout plane (31x31 cm2)
70 mm resolution achieved by strips centroid
 Analog readout required (used the APV25 chip)
COMPASS GEM Tracker
A tracker of twenty two 31 cm x 31 cm GEM chambers successfully operated for years
Rates as high as 2.5 MHz/cm2

## Slide 6: GEM Rate capability
Triple GEM
Poli Lener, PhD Thesis - Rome 2005
Maximum rate expected for SoLID GEM trackers
~ 600 kHz/cm2
Multiplication stages shielded from each other: much reduced feed back
Slow moving positive ions localized to holes, away from the induction region
Most of the created electrons contribute to signal: can operate at low gains
Much higher rates compared to wire chambers: ~ 50 MHz/cm2

## Slide 7
Current proposal to instrument locations 5, 6, 7, and 8 with GEM:
30 GEM modules at each location: each module with a 10-degree angular width.
PVDIS GEM configuration
Plane | Z (cm) | RI (cm) | RO (cm) | Active area (m2)
5 | 150 | 55 | 115 | 2.7
6 | 190 | 65 | 140 | 4.0
7 | 290 | 105 | 200 | 7.6
8 | 310 | 115 | 215 | 8.6
total: |  |  |  | ~ 23
Outline of a GEM module
Largest GEM module size required: 100 cm x (20-38) cm

## Slide 8
Suggested readout scheme:
  a 2D readout optimized to get high accuracy in the f coordinate, lower but sufficient resolution in the r coordinate.
  each set of stripes parallel to one of the radial sides of the module: i.e. stripes at a 10-deg stereo angle to each other.
  strip pitch is 0.6 mm for locations 7 and 8;
  0.4 mm for locations 5 and 6.
  Issues:
    High capacitance in long readout strips; will signal to noise ration be an issue ?
    How well will the 10-deg stereo angle separated strip layers work ? – need to test with prototypes.
PVDIS GEM configuration

## Slide 9
For this readout scheme readout channel estimation
PVDIS GEM configuration
with 20% spares, we will need about 170 k channels.
Good news: cost of electronics going down – cost per channel for the RD51 SRS APV-25 based readout is estimated to be ~ $ 2.50 - $ 3.00 + R&D expenses to optimize electronics for SoLID needs.
The total cost of readout electronics can be less than $ 1 M
Plane | Z (cm) | RI (cm) | RO (cm) | # of channels
5 | 150 | 55 | 115 | 30 k
6 | 190 | 65 | 140 | 36 k
7 | 290 | 105 | 200 | 35 k
8 | 310 | 115 | 215 | 38 k
total: |  |  |  | 140 k

## Slide 10
Six locations instrumented with GEM:
PVDIS GEM modules can be re-arranged to make all chamber layers for SIDIS. – move the PVDIS modules closer to the axis so that they are next to each other
More than enough electronic channels from PVDIS setup.
The two configurations will work well with no need for new GEM or electronics fabrication.
SIDIS GEM configuration
Plane | Z (cm) | RI (cm) | RO (cm) | Activearea | # of channels
1 | 197 | 46 | 76 | 1.1 | 24 k
2 | 250 | 28 | 93 | 2.5 | 30 k
3 | 290 | 31 | 107 | 3.3 | 33 k
4 | 352 | 39 | 135 | 5.2 | 28 k
5 | 435 | 49 | 95 | 2.1 | 20 k
6 | 592 | 67 | 127 | 3.7 | 26 k
total: |  |  |  | ~18 | ~ 161 k
PVDIS
SIDIS

## Slide 11
large area GEM chamber challenges
SoLID needs GEM modules as large at 100 cm x 40 cm.
The biggest challenge used to be the non-availability of large area GEM foils.
Not a problem anymore: CERN shop can make foils as large as 200 cm x 50 cm now.
One problem may be the production capacity of the CERN shop: especially if a LHC related large GEM project gets underway.
Currently work going on to develop large GEM production capabilities in China.

## Slide 12
Major recent development at CERN PCB shop towards large GEM foils
Base material only ~ 45 cm wide roll.
Used a double mask technique for etching: hard to the two masks accurately: Max area limited to ~ 45 cm x 45 cm previously.
Bias top surface to – w.r.t chemical bath
Double Mask
Single Mask
Single Mask technique allows to make GEM foils as large as 200 cm x 50 cm

## Slide 13
Major recent development towards large GEM foils
Splicing GEM foils together: seam is only 2 mm wide
Performance of the rest of the GEM foil unaffected

## Slide 14
TOTEM T1 prototype made with single mask GEM foils (33 cm x 66 cm)
Large prototype GEM module for CMS: 99 cm x (22 – 45.5) cm
CMS prototype similar the the dimensions of largest SoLID chambers
STAR Forward GEM Tracker
  6 triple-GEM disks around beam
  IR~10.5 cm, OR~39 cm
  APV25 electronics
Large GEM chamber projects

## Slide 15: Jefferson Lab SBSTrackers – Front Tracker
Front Tracker
Geometry
x6
18 modules
In Italy
50 cm x 40 cm Modules are assembled to form
larger chambers with different sizes
Front tracker: Six 40 cm x 150 cm GEM layers
  INFN Funding: to be built in Italy
Polarimeter trackers: Eight 50 cm x 200 cm GEM layers
  to be built in Virginia
Total GEM coverage of SBS ~ 16 m2
76 modules
In Virginia
SBS GEM module construction currently going on in Italy and at UVa

## Slide 16: The currenty ongoing GEM chamber work
Stretching time ~ 1 – 2 hours
Foil Tension: T ~ 0.5 kg/cm
New 40 cm x 50 cm GEM stretcher at Uva
Fully equipped INFN prototype 40x50 cm2 GEM module for a beam test at DESY
A 100 cm x (25-40) cm prototype GEM module with 2D readout is currently under design at Uva. (EIC detector R&D effort)
Expect to construct and test this chamber this Fall - Can learn a great deal for SoLID GEMs from this:
  develop techniques for fabrication of large GEM modules.
  study readout schemes: 10 degree stereo angle
  Study issues of signal to noise ratio for long readout strips.

## Slide 17: GEM chamber electronics
The RD-51 Scalable Readout System provides a low-cost, common platform that can accommodate different readout chips.
Currently tested with APV25-S1 chip.
Drawback with the APV25 chip: may not be fast enough for SoLID
Need to work on finding a suitable chip for SoLID readout and incorporating it into SRS (?)
SRS system has the benefit of the large team effort backed by RD-51
RD-51 plans to commercialize the fabrication; there will be the possibility to get very large systems in the future.

## Slide 18
Uva group now owns an APV25 system (2000 chan.) from the RD-51 Scalable Readout System (SRS) development.

## Slide 19: SoLID GEM chamber Collaboration
University of Science and Technology of China (USTC)
Tsinghua University
Lanzhou University
China Institute of Atomic Energy (CIAE)
Institute of Modern physics
INFN-Rome
University of Virginia

## Slide 20: rough cost estimate
Item | Quantity | Unit cost | Total cost | Material only unit cost | Material only total cost
GEM foil | ~100 m2 | $2000/m2 | 0.2 M | $3000/m2 | 0.2 M
readout boards | 120 | $ 2500 | 0.3 M | $ 2500 | 0.3 M
chamber support frame | 120 | $ 1500 | 0.2 M | $ 1500 | 0.2 M
Supplies and tooling |  |  | 0.1 M |  | 0.1 M
FEE and DAQ | 200 k | $ 5.0 | 1.0 M | $ 3.0 | 0.6 M
cables, power, etc |  |  | 0.5 M |  | 0.5 M
Gas system |  |  | 0.1 M |  | 0.1 M
Labor: Technicians | 6 FTE-years | $ 100 k | 0.6 M |  | -
Labor: Grad students | 6 student-years | $ 50 k | 0.3 M |  | -
support structure and integration |  |  | ??? |  | ???
TOTAL: |  |  | ~ 3.3 M |  | ~ 2.0 M
With 33% contingency |  |  | ~4.4 M |  | ~2.7 M
R&D and prototyping expenses: ~ $ 300 k

## Slide 21: Bottom-line
With the strong effort led by RD-51 towards large GEM chambers and readout electronics:
  Costs are coming down
  Many technical challenges are overcome
But SoLID is a very large and very complicated GEM project from any standard; no one has ever done anything even close to this size.
We are now starting on SBS GEM construction, this is about 70% of the size of the SoLID GEM project.
So with:
  large GEM chamber construction for SBS in Rome and at UVa
  the infrastructure and expertise we gain from large GEM prototyping at UVa, and
  Expertise gained by our Chinese collaborators.
  we will be ready for the SoLID GEM construction in a couple of years.